CHANGING SETTING FROM F TO C

- REMOVE THE TOP LID OF THE CABINET AND THE STAINLESS STEEL PLATE DIRECTLY UNDERNEATH.

- LOCATE THE BLACK BUTTON JUST ON THE BACK SIDE OF THE HEADER ASSEMBLY (DUAL CABINETS HAVE 2 - ONE FOR EACH CONTROLLER.)

- WITH THE UNIT POWERED ON, THE BLACK BUTTON (FOR THE APPROPRIATE CONTROLLER) WILL NEED TO BE DEPRESSED AND HELD THROUGHOUT THE CHANGE.

- WHILE THE CONTROL IS READING THE SET TEMP/ CURRENT TEMP, DEPRESS AND HOLD THE SET BUTTON FOR 3-5 SECONDS, ENTERING THE PARAMETERS. YOU WILL SEE THE DISPLAY READ “ON” THE TOP LINE AND THE WORD “OFF” ON THE BOTTOM. THEN PRESS THE SET BUTTON 7 MORE TIMES UNTILL YOUR SEE THE DISPLAY READ “UNIT” ON THE TOP LINE AND THE LETTER “F” ON THE BOTTOM.

- PRESS THE DOWN ARROW ONCE, CHANGING THE “F” TO A “C”.

- PRESS AND HOLD THE SET BUTTON AGAIN FOR 3-5 SECONDS, EXITING THE PARAMETERS.

- NOW PRESS AND HOLD THE UP ARROW AND THE SET BUTTON AT THE SAME TIME FOR 3-5 SECONDS UNTILL THE DIGITS “JIC / INT” APPEAR ON THE DISPLAY, AND RELEASE.

- PRESS THE SET BUTTON ONCE AND MOVE TO THE PARAMETER READING “H-SC”. BY HOLDING THE DOWN ARROW DEPRESSED, LOWER THE NUMBER 800 DOWN TO 71.

- PRESS THE SET BUTTON ONCE AGAIN, CHANGING THE PARAMETER TO “L-SC”. BY PRESSING THE UP ARROW, RAISE THIS NUMBER FROM 0 TO 32. THEN PRESS AND HOLD THE SET BUTTON FOR 3-5 SECONDS GETTING OUT OF THE PARAMETERS.

- PRESS AND HOLD THE SET BUTTON FOR 3-5 SECONDS AGAIN. ONCE IN THE PARAMETERS, PRESS THE SET BUTTON 6 TIMES UNTILL THE DISPLAY READS “AL “ON THE TOP LINE. USING THE UP ARROW CHANGE THE NUMBER TO 5. ONCE COMPLETED, PRESS AND HOLD THE SET BUTTON FOR 3-5 SECONDS EXITING THE PARAMETERS.  YOUR CABINET IS NOW CONVERTED TO CELCIUS.

- ADJUST THE TEMP TO YOUR NEW DESIRED SET POINT.